Performance Test Report

| Manufacturer |  | H&abyz Co., Ltd. |
| --- | --- | --- |
| Product name |  | Digital X-ray Imaging System |
| Model Name |  | HnX-R1 |
| Ref no. | (G series) | G1417CW, G1417CWP, G1417FW, G1417FWP, G1717CW, G1717CWP, G1717FW, G1717FWP |
|  | (2/3G series) | A1717MCW, A1417MCW, F1417MCW |
|  | Console software | HnVUE |

|  | Write | Review | Approval |
| --- | --- | --- | --- |
| Name | Han, Jimin | Kim, Myeongseop | Hur, Namkyu |
| Dept. | R&D Team | R&D Team Leader | Quality Management Representative |
| Date | 2025.07.18 | 2025.07.18 | 2025.07.18 |
| Signature |  |  |  |

.

|  |  |  | 목 차 |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |

0.	제·개정이력Revision History	3

1.	개요 Document Overview	4

2.	시험 정보Device information	4

2.1	시험품목	4

2.2	시험환경	4

3.	Test data 요약	5

4.	Test contents 시험 내용	6

4.1	의료용 영상의 획득, 저장, 전송 등의 기능	6

4.2	촬영된 환자의 신상정보, 촬영 오더 정보, 촬영 영상 등의 저장, 조회	7

4.3	영상의 확대, 축소, 회전, 이동, 밝기 조절 등의 영상 조작	8

4.4	영상의 레이아웃 표시 선택, Thumbnail 영상포시, 부분 확대 등의 기능	8

4.5	DICOM 3.0규격지원 영상 포맷 및 전송, 저장기능	9

4.6	사용 권한을 부여받은 사용자 ID	10

4.7	영상삭제 권한 승인	10

4.8	PACS 서버 오류 시 해결방법 제공	11

4.9	오류 발생 시 안내 메시지	12

4.10	촬영된 영상의 환자정보 확인	12

4.11	PACS와 통신 연결	13

5.	결론	14

6.	기타 첨부서류	14

Revision History

| Rev. No. | Issue date | Note |
| --- | --- | --- |
| 0 | 2025.07.18 | First establishment |

Document Overview

This test report is a test report to confirm the performance of the Console software(HnVUE) of the Digital X-ray Imaging System (HnX-R1).

Device information

Test item

- Product name: Digital X-ray Imaging System

- Model name: HnX-R1

- Test product: HnVUE

| Console Software | Ref no. | Version |
| --- | --- | --- |
|  | HnVUE | 1.0.0.37 |

Test environment

- Test date: 2025.07.14~2025.07.18

- Test site:  H&abyz Co., Ltd. (2-dong, 41-16 Cheoinseong-ro, Namsa-myeon, Cheoin-gu, Yongin-si Gyeonggi-do, Korea Rep.)

Summary of Test Results

| No | Test item | Test criteria | Test method | Test result |
| --- | --- | --- | --- | --- |
| 1 | Performance | Capable of acquiring, storing, and transmitting medical images. | Verify the acquisition, storage, and transmission functions of medical images. | PASS |
| 2 | Performance | Storage and retrieval of patient information, imaging order information, and captured images are possible. | Check each piece of information in the captured images. | PASS |
| 3 | Performance | Image manipulations include zooming in and out, rotating, moving, and adjusting brightness. | Check each operation in the captured images. | PASS |
| 4 | Performance | Functions include selecting image layouts, displaying thumbnails, and partial magnification. | Check each function in the captured images. | PASS |
| 5 | Performance | DICOM 3.0 standard-compliant image formats and transmission and storage capabilities are available. | 1. Click Close Study or Send to confirm DICOM transmission by checking the blue light in the PACS status window. 2. Click Print to confirm DICOM printing. | PASS |
| 6 | Security | Logging in with an authorized user ID is required to operate the imaging software, and data and information are protected against unauthorized access by anyone other than authorized users. | When using the Console software, ensure that access is only possible with the assigned user ID. | PASS |
| 7 | Security | Image deletion is performed after an authorization process to prevent accidental or intentional deletion. | Verify that the authorization process is in place when deleting images. | PASS |
| 8 | Security | Procedures and solutions are provided when the PACS server is down and image transmission is impossible. | Verify that PACS server downtime handling and resolution methods are provided. | PASS |
| 9 | Security | In the event of a hardware or software error, an on-screen notification message is provided. | Verify that on-screen guidance messages are provided in the event of hardware or software errors. | PASS |
| 10 | Security | Patient information from captured images can be viewed. | Verify patient information in captured images. | PASS |
| 11 | Security | The device can be used through a communication connection with the PACS. | Verify the connection between the PACS and the detector. | PASS |

Test contents

Functions for acquiring, storing, and transmitting medical images

| No. | Test items | Test criteria | Test method | Test result |
| --- | --- | --- | --- | --- |
| 1 | Performance | Capable of acquiring, storing, and transmitting medical images. | Check the capabilities of acquiring, storing, and transmitting medical images. | PASS |

* Test data

Storage and retrieval of patient personal information, shooting order information, and shooting videos

| No. | Test items | Test criteria | Test method | Test result |
| --- | --- | --- | --- | --- |
| 2 | Performance | Patient information, imaging order information, and captured images can be stored and retrieved. | Check each piece of information in the captured images. | PASS |

* Test data

Image manipulation such as zooming in, out, rotating, moving, and adjusting the brightness of the image

| No. | Test items | Test criteria | Test method | Test result |
| --- | --- | --- | --- | --- |
| 3 | Performance | Video manipulations such as zooming in, zooming out, rotating, moving, and adjusting brightness are possible. | Check each manipulation of the captured video. | PASS |

* Test data

Features such as selecting the layout of the video, displaying thumbnails, and zooming in on a part of the video

| No. | Test items | Test criteria | Test method | Test result |
| --- | --- | --- | --- | --- |
| 4 | Performance | Features include selecting the video layout, thumbnail display, and partial zoom. | Check each feature of the captured video. | PASS |

* Test data

DICOM 3.0 standard support image formats and transmission and storage functions

| No. | Test items | Test criteria | Test method | Test result |
| --- | --- | --- | --- | --- |
| 5 | Performance | DICOM 3.0 standard-compliant image formats, transmission, and storage capabilities are available. | 1. Click Close Study or Send to confirm DICOM transmission by checking the blue light in the PACS status window. 2. Click Print to confirm DICOM printing. | PASS |

* Test data

User ID granted permission

| No. | Test items | Test criteria | Test method | Test result |
| --- | --- | --- | --- | --- |
| 6 | Security | To operate the video software, you must log in with an authorized user ID. Information and data are protected, preventing access by anyone other than authorized users. | When using the console software, ensure that access is restricted to authorized user IDs. | PASS |

* Test data

Approval of video deletion permission

| No. | Test items | Test criteria | Test method | Test result |
| --- | --- | --- | --- | --- |
| 7 | Security | To prevent accidental or intentional deletion of videos, deletions are processed through a permission approval process. | Verify that the approval process is running when deleting a video. | PASS |

* Test data

Provides solutions for PACS server errors

| No. | Test items | Test criteria | Test method | Test result |
| --- | --- | --- | --- | --- |
| 8 | Security | Provides solutions and troubleshooting methods when the PACS server is down and image transmission is impossible. | Check if solutions and troubleshooting methods are available when the PACS server is down. | PASS |

* Test data

Guidance message when an error occurs

| No. | Test items | Test criteria | Test method | Test result |
| --- | --- | --- | --- | --- |
| 9 | Security | If a hardware or software error occurs, a notification message is displayed on the screen. | Check whether a notification message is displayed on the screen when a hardware or software error occurs. | PASS |

* Test data

Confirming patient information in recorded images

| No. | Test items | Test criteria | Test method | Test result |
| --- | --- | --- | --- | --- |
| 10 | Security | You can check patient information from the recorded images. | Check patient information from the recorded images. | PASS |

* Test data

Communication with PACS

| No. | Test items | Test criteria | Test method | Test result |
| --- | --- | --- | --- | --- |
| 11 | Security | The device can be used through a communication connection with the PACS. | Check the connection between the PACS and the detector. | PASS |

* Test data

Conclusion

The test data confirmed that the medical imaging device (model name: HnX-R1, HnVUE) performed without any issues.

Other Attachments

[Attachment 1.] ISO 13485 Certificate

[Attachment 1.] ISO 13485 Certificate